From:

FORMED Solutions INC

92, Coralland, Haggatt Hall, ST. Michael Barbados

Barbados 15/12/2024

# Quotation request

Hi,

We request from you the quotation for the below mentioned equipment.

- Autopsy Saw with vacuum unit of at least 7 litres container capacity, air volume of 32 litres/second, filtration with a 99.9% at 0.6-micron, voltage rating 120V/1 phase/60Hz.

Handpiece

DC permanent magnet Speed:40VoltDc/0.75 Amp Typical load: 1.5 Amps Speed:1200/min

Construction: Aluminium body bearing sealed mechanism protected to IEC IP67 standard (waterproof)

Accessories Saw blades, service kit, vacuum hood, micro filter, vacuum pipe.

- Autopsy Table

  - Central support column with an accessor door

  - Rotation to left or right at an angle of 180 degree.

  - Inclusive of a basin W/D/H within range of 400mmx500mmx200mm

  - Outlet strainer

  - Knee operated mixing tap for cold and warm water.

  - Shower hose with hand sprinkler to be included.

  - Splash proof electric outlet.

  - Buttons for vertical height adjustment

  - Preinstalled electric and water connections

  - Voltage 120V/1Ph/60Hz

  - Carrying capacity of 200kg or 440lbs

- Electronic Organ scale

  - Stainless steel construction

  - Clear LCD display

  - Weight must be switched from kg to lbs.

  - Automatic zero setting, tare function and hold function.

  - Removable stainless- steel bowl

  - Operation by mains, 120V, 60Hz and rechargeable batteries

  - Accuracy of 1 gram

  - Documentation: Operator’s manual.

- Autopsy Transport Trolley

  - Heavy duty, rugged, torsion-resistant, lightweight, steel construction with antibacterial powder coat.

  - Capable of manual operation by single operator

  - Dimension of no less 84’’ in length and 30’’ in width

  - Adjustable height range with 12’’ to 65’’

  - Carrying capacity of not less than 500lbs

  - Hydraulic foot controls mounted on the same side as steering handle

  - 5’’ maintenance free solid rubber wheel for steering and noiseless movement

  - Locking mechanism built in to keep trolley in place

  - Unit must be power washable and have removable handle

  - Operators’ manual

- Fresh Air Hood

  - Stainless steel construction

  - Integrated lamps of no less than 58Watt for provision of regular light, colour temperature of approximate 5500 Kelvin

  - Dimension W/D/H: 1000x2400x300mm

  - Air Capacity 750m3/h

  - Trajectory range: 2.0m

  - Loss of pressure: approximate 40 pa

  - Adjustable air conduct

  - Exhaust duct diameter 2xd=200mm

  - Voltage 120V/1ph/60Hz

- Autopsy Instrument Kit

  - 1 piece Scalpel with metal handle, cutting length 37mm

  - 1 piece Cartilage knife with metal handle, cutting length 55mm

  - 1 piece bone shears 225mm

  - 2-piece dissecting forceps

  - 1 piece bone saw 350mm, inclusive of spare blade

  - 1 piece metal mallet, 235mm, 300 grams

  - 2 piece Osteotome, 130mm

  - 1 piece Rachiotome, 220mm

  - 2 piece Blow pipe, straight, double, 104mm

  - 2 piece Retractors, double, 65mm

  - 2 piece Probes 200/1.5mm

  - 1 piece Surgical needles (12 pieces)

  - 1 piece folding ruler 300mm

  - 1 piece measuring glass 100cm

  - Stainless steel sterilizing Box/container